#FunTimesWithTheTA 
Galvanic Skin Response

Adapted from: https://www.element14.com/community/groups/pumping-station-one/blog/2011/05/08/ouch-sensing-galvanic-skin-response-gsr by Drew Fustini

Introduction

In this installment of #FunTimesWithTheTA we will build a circuit for measuring Galvanic Skin Response (GSR). With GSR, we will be able to collect measure changes in skin conductance, which is an indirect measurement of sympathetic nervous response (or “flight-or-fight” response). By doing so, we can sense when someone is in an excitatory state elicited by some sudden change in emotion.

SAFETY NOTICE: This circuit measures the electrical activity of a small section of the human body. Though the risk of injury is miniscule, every effort to ensure safety must be made. As a result, when connecting the circuit to a person, the circuit must not be connected to any component or instrument connected to mains power supply. Instead, please power your circuit with a small CR2032 coin cell battery. These batteries provide very little risk of shock in the case of a short circuit due to their small output current capability.

Figure 2: Circuit diagram of GSR Circuit.

Table of Contents

| Section | Page |
| --- | --- |
| Introduction | 1 |
| Objectives | 1 |
| GSR Schematic | 1 |
| Part #1: Voltage Reference Generator (Vref) | 3 |
| Part #2: Transducer: Electrode Contacts and Voltage Divider | 5 |
| Part #3: Active Bandpass Filter | 6 |
| Part #4: 2nd Order Low-Pass Filter | 7 |
| Part #5: Bypass Capacitor | 8 |
| Revision History | 9 |

Part #1: Voltage Reference Generator (Vref)

Figure 3: Voltage Reference Generator using a voltage divider and unity gain buffer.

Single Supply Op Amps

In this circuit, we are utilizing single supply op amps. These op amps are powered by a single voltage source, like a small battery, and ground. This is in contrast to op amps that are powered with a dual voltage (for example +/- 15V) like you may have used in the past. The specific op amp we are using is called the MCP6002. It is a general purpose single supply, rail-to-rail op amp from Microchip Technology. The specific name, MCP6002, isn’t important. MCP6002 is just a product code from the manufacturer.

Figure 4. Graphic displaying the utility of the voltage reference generator in ensuring that AC signals stay within the voltage supplies.

For using AC signals, single supply op amps present an interesting challenge. If you remember from your basic filter characteristics, high-pass filters block DC voltage. This means, if we put a 1Vpp sine wave with a 1V DC offset into a high pass filter the 1V DC offset gets filtered (because high pass filters block low frequencies and DC is a low frequency i.e. 0Hz) resulting in a 1Vpp sine wave that fluctuates between -500mV and +500mV centered around 0V. Our single supply op amp does not have a negative supply voltage, meaning that we could not create a sine wave that goes below 0V. As a result, we have to force the middle point of the sine wave to be a voltage between the supply rails, ground and our positive supply. We do that with a voltage reference or Vref generator.

A voltage reference generator can be made in made in a variety of ways. One of the most common ways is by using a voltage divider and a voltage buffer (unity gain amplifier) as show in Figure 3. Recall that the voltage at both terminals of the op amp are the same. Therefore, if we hold the non-inverting pin to a certain voltage within the supply rails and tie the inverting pin to the output of the op amp, the following equation must hold true:

Vout=VIN+ 						(1)

where VIN+ is the voltage at the non-inverting pin set by the voltage divider. Furthermore, recall that for a voltage divider (using the component labels in our schematic)

Vout,voltage divider=VIN(R9(R9+R8)) 			(2)

Most all parts of the circuit need this reference voltage, so please build this stage first.

Testing

- Power the op amps with 3V and GND from the benchtop power supply. Please see Part #5 for demonstrating the use of bypass capacitors.

- Assemble the circuit as shown in Figure x

- Compare the voltage at the non-inverting pin to the voltage at the output of the op amp. They should be the same.

Part #2: Transducer: Electrodes and Voltage Divider

Electrodes

Electrodes are components used to interface between an electrical circuit and some sort of conducting medium. In this case, we will be using electrodes to connect to a small portion of the skin and measure small changes in skin conductance (inverse of resistance). We will be using copper tape as electrodes. The specific copper tape we will be using has copper on both sides (3M, DigiKey Part Number: 3M1126C-1YD-ND), but the two sides are electrically isolated (there is no electrical connection between both sides of the tape). As a result, please connect both sides of the tape together by wrapping the tape fully around your finger and sticking one side to the other.

Voltage Divider

Figure 5: Voltage Divider Circuit for Measuring Change in Skin Conductance (inverse of resistance).

We will be creating a simple voltage divider circuit to measure changes in skin conductance. As you recall from equation 2, the output of the voltage divider changes with the ratio of the two resistor values. If one of the resistor values changes (the skin) the output voltage of the voltage divider circuit will change accordingly. You may be able measure GSR fine this way. Otherwise, we will be amplifying the circuit a bit more to see the response a bit more clearly.

SAFETY NOTICE: This circuit measures the electrical activity of a small section of the human body. Though the risk of injury is miniscule, every effort to ensure safety must be made. As a result, when connecting the circuit to a person, the circuit must not be connected to any component or instrument connected to mains power supply. Instead, please power your circuit with a small CR2032 coin cell battery. These batteries provide very little risk of shock in the case of a short circuit due to their small output current capability.

Testing Instructions

- Obtain two pieces of copper tape that are long enough to wrap around your finger

- Solder a wire to each of the two pieces of copper tape

- Attach the copper tape to the voltage divider circuit

- Power your circuit with a small CR2032 coin cell battery

- Measure the output of the voltage divider circuit using an isolated measurement instrument. (You will be provided with a battery powered oscilloscope or microcontroller that can be used to view the signal).

Part #3: Active Bandpass Filter

Figure 6: An active bandpass filter to attenuate noise and selectively amplify the AC galvanic skin response.

A band-pass filter is a useful filter configuration that allows us to combine the effects of low-pass and high-pass filters resulting in only a “band” of frequencies in the passband. We will use a bandpass filter to block DC signals to ensure that only the AC portion of our signal is being amplified and for filtering high frequency noise such as 60 Hz. The band-pass filter has two cut-off frequencies set by the low-pass and high-pass filter. Recall from equation 3 that the cut-off frequency of low-pass and high-pass filters are determined by

fc=12πRC			(3)

The low-pass filter determines the higher cut-off frequency. The high-pass filter determines the lower cut-off frequency.

Notice that our bandpass filter has a gain of one in the passband:

Gain=R1R2					(4)

This is intentional. We will amplify the signal in the next stage.

Testing

- Power the op amps with 3V and GND from the benchtop power supply. Please see Part #5 for demonstrating the use of bypass capacitors.

- Create a 1 Hz, 2 Vpp sine wave from your function generator. Use High-Z mode.

- Input the sine wave into your bandpass filter.

- You should have an output signal of about 1.76 Vpp, just slightly smaller than the input voltage. This corresponds to a gain of about -1 dB.

Try It!

- Try changing the frequency of your input sine wave to 60Hz.

- Using, High-Z mode, keep the signal amplitude at 2 Vpp.

- You should have an output signal of about 160 mVpp.

- This corresponds to about -22dB attenuation.

Part #4: 2nd Order Low-Pass Filter

2nd Order Low-Pass Filters

Figure 7: A second order low-pass filter to filter out 60Hz noise. This filter is more helpful than a 1st order filter since the transition band is steeper (-40dB/decade).

You are already familiar with 1st order high-pass and low-pass filters. These require only a single resistor and capacitor. 1st order filters are really helpful because they are simple to construct and require minimal number of components. However, sometimes you may need better filter performance. 2nd order filters are perfect for as these have a steeper roll-off in the transition band (-40dB/decade) compared to 1st order filters (-20dB/decade). We will use our 2nd order low-pass filter specifically to filter out 60Hz noise. The human body, by nature, acts like a giant antenna and picks up stray electromagnetic radiation, especially 60Hz noise. Since we electrodes connected to the skin to measure skin conductance, we actually get quite a bit of 60Hz noise coupling through our electrodes to the rest of the circuit through our bodies. The 2nd order low-pass filter helps to attenuate the noise. There a few equations that we could discuss to characterize the 2nd order low-pass filter, but we chose the simpler configuration. If both resistors and capacitors are the same values, the cutoff frequency of the filter is the same as you have been using (Equation 3). Now we get much more attenuation at 60Hz compared to if we simply used a 1st order filter.

Testing

- Power the op amps with 3V and GND from the benchtop power supply. Please see Part #5 for demonstrating the use of bypass capacitors.

- Use High-Z mode. Create a 60 Hz, 2.5 Vpp sine wave from your function generator.

- Input the sine wave into your 2nd order low-pass filter.

- You should have an output signal of about 62 mVpp. This corresponds to a gain of about -32 dB.

- Try changing the frequency of your input sine wave to 1Hz.

- Use High-Z mode. Keep the signal amplitude at 2.5 Vpp.

- You should have an output signal of about 2.5 Vpp showing no attenuation in the passband.

Part #5: Bypass Capacitors

Bypass Capacitors

Figure 8: Bypass capacitors placed as close as possible to the power supply pin of each op amp in the circuit. These bypass capacitors help reduce power supply noise as well as keep the operation of the circuit stable.

Image that instead of providing a constant DC supply, our batteries or benchtop power supply actually have a small AC noise on top of the DC voltage. Imagine that sometimes, this AC noise can become rather large. If this AC noise is too large and drops below the operating voltage of our op amps, the circuit can momentarily turn-off. As you imagine, this could potentially be pretty bad. It would be really awkward if you were using your cellphone and all of a sudden, it would suddenly turn-off. This is the unfortunate reality of real voltage sources. Real voltages sources cannot provide a truly steady voltage level, but droop and spike with changes in load (the circuit the DC supply is powering). We can minimize the effects of these momentary changes in power supply voltages with capacitors. Recall that capacitors resist instantaneous changes in voltage. Also recall that capacitors actually store charge. If we connect a capacitor directly across the terminals of the battery, the capacitor will charge up to the voltage level of the battery. Since capacitors resist changes in voltage, the capacitor reduces AC noise on our DC voltage and help keep the operating voltage of our circuit constant. In this configuration, capacitors are referred to as bypass capacitors. It is important to use bypass capacitors for each one of your op amps and to place these capacitors as close as possible to the power supply pin of your op amp. Bypass capacitors also help to stabilize the operating

Please see the following video to see the effect of using a bypass capacitor in a practical circuit. <https://youtu.be/fr-S1TsHXvk>

Revision History

| Revision Code | Revision Date | Description |
| --- | --- | --- |
| A | 10/31/2017 | - Initial document for building the Galvanic Skin Response Circuit for #FunTimesWithTheTA 2017. Circuit based on design by Drew Fustini from Pumping Station One and Element14.com <https://www.element14.com/community/groups/pumping-station-one/blog/2011/05/08/ouch-sensing-galvanic-skin-response-gsr> |
| A.2 | 11/14/2017 | - Changed equation 4 to reflect component designators shown in figure 6 |
| A.3 | 8/2/2018 | - Minor revisions fixing typos and similar errors |